# OpenBCI EEG Viewer – Technical Report

Document: Cogwear Demo EEG Viewer

Version: Multi-headband real-time viewer with synchrony analysis

# 1. Overview

The EEG Viewer is a real-time visualization and analysis application for OpenBCI Cyton headbands. It supports 1–4 headbands simultaneously (4–16 channels), with preprocessing, frequency-band decomposition, and inter-subject EEG synchrony (Phase Locking Value) when multiple headbands are connected.

# 2. Hardware Support

| Feature | Specification |
| --- | --- |
| Headbands | 1–4 OpenBCI Cyton headbands |
| Channels per headband | 4 (AF7, FP1, FP2, AF8) |
| Total channels | 4, 8, 12, or 16 |
| Sample rate | 250 Hz |
| Connection | USB dongle per headband → COM port (e.g. COM3, COM4) |
| Radio | Each headband must use a different radio channel (see RADIO_SETUP.md) |

# 3. Usage

| Command | Description |
| --- | --- |
| python view_raw_signal.py COM3 | Single headband |
| python view_raw_signal.py COM3 COM4 | Two headbands |
| python view_raw_signal.py COM3 COM4 COM5 COM6 | Four headbands (default when no ports specified) |

# 4. Display Windows

Raw EEG – Preprocessed signal (0.3–50 Hz bandpass, notch, ICA, z-score) in 5-second rolling window

Theta (4–8 Hz) – Band-filtered theta activity

Alpha (8–13 Hz) – Band-filtered alpha activity

Beta (13–30 Hz) – Band-filtered beta activity

Gamma (30–60 Hz) – Band-filtered gamma activity

EEG Synchrony (PLV) – Pairwise Phase Locking Value in alpha band (8–13 Hz); shown only when 2+ headbands are connected

Plots are arranged in a grid (2×2 for 4 channels, 2×4 for 8, 4×4 for 16). Packet loss and flat-signal indicators appear at the bottom of the Raw EEG window.

# 5. Signal Processing Pipeline

## 5.1 Preprocessing (applied to all channels)

| Step | Setting | Value |
| --- | --- | --- |
| Bandpass | Low / High cutoff | 0.3 Hz / 50 Hz (4th-order Butterworth) |
| Notch | Frequency / Q | 50 Hz (or 60 Hz) / Q=30 |
| Common Average Reference (CAR) | Per headband | Subtract mean of 4 channels from each channel within that headband |
| ICA | FastICA, kurtosis-based rejection | Reject components with ¦z¦ > 1.96; refit every 2.5 s |

## 5.2 Display-Only Processing

These steps affect visualization only, not stored or exported data:

Median filter (5-point) to reduce spikes

Outlier rejection: replace values beyond ±2 z-score with local median

Moving average smoothing (15-point)

Display clipping at ±2.5 z-score

Y-axis auto-range

# 6. EEG Synchrony (Multi-Headband)

When 2 or more headbands are connected: Phase Locking Value (PLV) is computed in the alpha band (8–13 Hz) for all pairwise combinations. Each headband's signal is the mean of its 4 channels. PLV ranges from 0 (no phase locking) to 1 (strong phase locking). Bar chart updates in real time (~30 Hz).

# 7. Diagnostics and Monitoring

Packet loss – Per-headband count and percentage

Flat-signal indicator – "⚠flat" when a headband's signal is much flatter than others

Cross-talk check – Warns if two headbands show correlation > 0.92 (possible same radio channel)

Radio channel reminder – Startup message when 2+ headbands; see RADIO_SETUP.md

# 8. Configuration Summary

| LOW_CUTOFF | 0.3 Hz |
| --- | --- |
| HIGH_CUTOFF | 50 Hz |
| NOTCH_HZ | 50 (or 60 for Americas) |
| NOTCH_Q | 30 |
| USE_CAR | True |
| ICA_DISABLED | False |
| Z_THRESHOLD | 1.96 |
| ICA_UPDATE_INTERVAL_MS | 2500 |
| PLOT_REFRESH_MS | 33 (~30 Hz) |
| DISPLAY_SMOOTH | 15 |
| OUTLIER_REJECT_Z | 2.0 |
| DISPLAY_CLIP_Z | 2.5 |
| DISPLAY_MEDIAN_FILTER | 5 |
| SYNC_BAND (PLV) | 8–13 Hz (alpha) |

# 9. Dependencies

Python 3, pyqtgraph, PyQt5, numpy, scipy, scikit-learn, pyserial

# 10. File Structure

view_raw_signal.py – Main viewer application

lib/open_bci_v3.py – OpenBCI serial protocol and board interface

RADIO_SETUP.md – Radio channel configuration guide